Amy Henrickson,
Saeed Mortezazadeh, and
Borries Demeler
ABSTRACT
Accurate characterization of adeno-associated virus (AAV) formulations is an essential requirement for several reasons. It has been shown that dose-dependent capsid antigen presentation leads to reduced transduction and proteasomal degradation of viral capsids, and ultimately to therapeutic failure. It is therefore essential to accurately characterize nucleic acid cargo loading of AAV virions, and to develop methodology to a) efficiently separate empty, aggregated and partially filled capsids from filled virions, and b) to be able to quantify a gene therapy drug formulation for the precise amount of properly filled capsids compared to the other species. Challenges encountered in the characterization of AAV formulations relate to the accuracy and resolution with which this quantification can be made, the sensitivity with which the measurement can be made, largely determining sample requirements, and the throughput and cost of the analysis. Multi-wavelength analytical ultracentrifugation (MW-AUC) offers two highly precise characterization tools to identify the amount of each capsid cargo load, as well as the presence of contaminants such as protein aggregates or free nucleic acid. Both methods resolve different species with high fidelity. The first method is performed in sedimentation velocity mode (SV), providing separation based on mass, size, anisotropy and density, and excels at identification of contaminants. The second method is performed in analytical buoyant density equilibrium (ABDE) mode, using a density gradient material, separating molecules based on their density alone. The ABDE method offers higher throughput and sensitivity, but does not differentiate based on size. Multi-wavelength decomposition of the hydrodynamic signal from both methods additionally provides an orthogonal, optical characterization dimension, separating signals from protein and DNA, further improving the resolution of this approach by identifying the fraction of each bio-polymer in each species. Together, these methods provide an unambiguous, first-principle identification of each virion.
MW-AUC provides high resolution solution composition information for mixtures heterogeneous in mass, globularity/anisotropy, density, size, and spectral properties. We show here optimizations of this method for the characterization of formulations of AAV virions that differ in nucleic acid cargo and can contain contaminants. Our optimizations focus on the following properties:
Sensitivity (enhancing accurate detection at low concentrations)
Resolution (identifying and quantifying individual virion populations)
Reducing requirements for sample volume and titer, as well as concentration
Eliminating systematic noise contributions to the observed signal
Enhancing signal-to-noise ratio by reducing remaining stochastic noise
Quantifying and separating protein and DNA signals by applying MW-AUC
Maximizing throughput (up to 8 samples/experiment in a 4 hole rotor)
Automated workflows with UltraScan GMP data acquisition and analysis

---
OPTIMIZATIONS
Raw Data Processing:
SV and ABDE experiments are optimized by collecting data in intensity mode. This eliminates the amplification of stochastic noise by a factor of √2 when a reference scan is subtracted from the sample scan [1]. Furthermore, the reference channel can be used for a second sample without adding to the scanning time, doubling throughput. For SV experiments, time- and radially- invariant noise contributions (TI/RI) are numerically removed in UltraScan [2]. For ABDE experiments, all data of interest are time invariant, and only the last scan is considered to be at equilibrium and analyzed. TI and RI noise contributions are managed by the reference scan module in UltraScan, which converts intensity data to pseudo-absorbance using the fingerprint of the photomultiplier tube in the Optima AUC as a function of wavelength to extract the TI background signal from experimental data without the addition of stochastic noise. This process is implemented in the UltraScan software to increase the signal-to-noise ratio and to improve data quality, and to double the capacity of the instrument. The result of the reference scan processing is illustrated in Fig. 1 below:
Temperature considerations:
For AAV SV experiments performed at 20ºC it was shown that in low-density buffers a convection effect is likely [3]. To avoid this issue, the temperature can be lowered to 4ºC to reduce the density gradient effect. At 4ºC, the density is nearly constant with small variations in temperature, which will avoid this problem. Furthermore, at 4ºC the density and viscosity of the buffer is highest, reducing the sedimentation speed of the relatively fast sedimenting AAV particles.
Improving Signal-to-Noise ratios:
The number of scans collected in a SV experiment is proportional to the measured signal. To maximize the number of scans, it helps to reduce the number of wavelengths collected. Increasing the density and viscosity by changing the temperature (see above) slows down sedimentation and increases the number of scans collected.Rotor speed optimization is especially important at low rotor speeds, where synchronization of the rotor hole position with the flash lamp frequency is essential, so the lamp does not have to wait for the rotor to be in the right position. The shortest scan times are achieved at 14.6 krpm on our instruments, with scan times of 8 seconds/channel.
Figure 1: Noise correction of an ABDE AAV experiment. Raw ABDE intensity data collected on the Optima AUC (only a single wavelength is shown) is shown in the left panel. The center panel shows an uncorrected pseudo-absorbance dataset. A baseline containing time-invariant noise contributions is clearly visible. The right panel shows the same data after processing the data with the UltraScan reference scan module, resulting in a mostly flat baseline.

---
AAV Characterization by Multi-wavelength
Analytical Ultracentrifugation
OPTIMIZATION (cont.)
Sample Size and Concentration:
AAV samples are relatively costly to prepare, so reduction of sample requirements for quality control is an important factor. While MW-AUC SV experiments require about 0.46 ml of sample at concentrations of around 0.6-1.0 OD at their peak absorbance wavelength, ABDE experiments require significantly less material. In a buoyant density experiment, absorbance requirements are significantly reduced, since all the sample will concentrate in a sharp peak at the isopycnic point in the gradient, with the sample concentration everywhere else being close to zero. The half-width of an AAV peak is approximately 0.02 cm. Therefore, an equivalent absorbance response in a 1.2 cm solution column can be achieved with ~ 30 fold lower concentration.
Density Gradient Optimization:
The resolution of the ABDE method can be tuned to the density of the particles to be separated. By adjusting rotor speed, the steepness of the gradient can be modulated, which is shown in Fig. 2.
Figure 2: Density gradient shape. The higher the rotor speed, the steeper the gradient will become (red: 60 krpm, yellow: 50 krpm, green: 30 krpm). Since the gradient shape is proportional to the density of the gradient, the rotor speed can be tuned to maximally resolve the target species based on their density. Furthermore, the concentration of the gradient material can be adjusted to shift each curve up or down, depending on the needed range. For resolving a maximum density range, it is helpful to use the highest speed possible.
Multi-Wavelength AUC:
AAV samples are typically scanned from 240-290 nm for samples containing higher protein-DNA ratios (27.6º orthogonality), and 210-240 nm with lower ratios (17.8º orthogonality). Since extinction is higher at shorter wavelength, lower concentrations can be measured with high signal below 245 nm. Spectral overlays are shown in Fig. 2:
Figure 3: Protein and DNA spectra overlays for two different wavelength ranges. Protein and DNA contributions in AAV samples can be spectrally deconvoluted because the spectra are not linearly dependent. The orthogonality can be measured by the angle between the two spectra.

---
METHODS
Multi-wavelength Analytical Ultracentrifugation:
All data are acquired with an Optima AUC (Beckman Coulter), using the UltraScan-III GMP data acquisition system [4]. Data analysis is performed with UltraScan-III (v. 6411) [2]. Calculations were performed on an UltraScan-in-a-Box server (AUC Solutions). All SV experiments are performed in standard PBS buffer, using epon centerpieces (Beckman Coulter). A single cell is loaded with two samples (460μl at ~0.6-0.8 OD260/280), and measured at 4.0ºC and 14.6 krpm in intensity mode. ABDE experiments are performed at 60 krpm using up to four titanium centerpieces (Nanolytics) and spun to equilibrium (~400 minutes). 10-20 μl of up to eight AAV samples at ~0.6-0.8 OD260/280 are added to 1x ClearDens buffer (AUC Solutions) to achieve a 1.2 cm solution column, and measured in intensity mode. SV experiments are scanned with 2 nm increments, ABDE experiments are measured with 1 nm increments since there are no restrictions on scanning time. The analysis workflow is illustrated in Fig. 4:
AAV 3D Multi-wavelength s-distributions
Pseudo-absorbance 4D Multi-wavelength data
(a single scan is shown)
LDNA Lprotein
NNLS Decomposition of individual wavelength scans at each radial position (cyan line above) into protein and DNA basis spectra results in traditional 2D datasets for protein and DNA signals alone
C_MW~=~ color red {x_a} left [ matrix { %epsilon_a1 ## %epsilon_a2 ## ... ## %epsilon_ai } right ]_{r,t} ~+~ color blue {x_b} left [ matrix { %epsilon_b1 ## %epsilon_b2 ## ... ## %epsilon_bi } right ]_{r,t}
NNLS
2D DNA data 2D Protein data
Figure 4: MW-AUC analysis workflow. 4D SV raw data scans are either fitted to s-value distributions or deconvoluted into protein and DNA contributions and analyzed separately. For deconvolutions, the wavelength scans at each radial position of each scan are deconvoluted, generating a single scalar Xa (DNA) or Xb (protein) representing a single point in a 2-dimensional traditional velocity dataset which can be processed with traditional methods. In ABDE experiments, the data at equilibrium are time-invariant, requiring only the last scan to be analyzed.

---
Department of Chemistry and Biochemistry
University of Lethbridge
Lethbridge, Alberta, Canada T1K 3M4
Results:
Sedimentation Velocity Experiments:
MW-AUC SV experiments of AAV samples offer access to a large range of particle sizes and molar masses, and provide excellent resolution. This means even small molecular weight contaminants as well as large aggregates can be unambiguously identified. In contrast to traditional 260/280 nm analysis, the relative amount of each species can be more accurately determined with MW-AUC, and the nature of a contaminant identified. Fig. 5 shows a SV analysis performed with the enhanced van Holde – Weischet method [5] for an AAV sample containing a small molecular weight nucleic acid contaminant (about 15% of the total DNA signal), a mostly empty capsid that contains a small amount of DNA (about 90% of capsid material), and about 10% of filled capsid:
Figure 5: MW-AUC analysis of an AAV sample. Left panel: 280 nm (cyan) and 260 nm (blue) analysis, showing the correct species, but failing to identify the biopolymer type of the ~20S contaminant, and to quantify the correct ratios of filled and partially filled AAV virions. Right panel: MW-AUC analysis of the same data analyzed in the left panel. Here, the 20S contaminant is clearly identified as containing only DNA, and the presumed empty capsid at around 65S is clearly co-migrating with a small amount of DNA, making it a partially filled species. The amount of filled capsid sedimenting at ~100S is overestimated by the 260/280 analysis, as only the protein signal is needed to derive correct ratios of empty, partial and filled species.

---
Analytical Buoyant Density Equilibrium Experiments (ABDE):
ABDE experiments provide high resolution information about the degree of DNA cargo loading of AAV virions. ABDE experiments distinguish based on the buoyant density of a solute inside a density gradient, and produces sharp peaks for each species which are easily quantified by a peak fitting algorithm. Compared to SV experiments, ABDE experiments have approximately 30 fold lower sample requirements, and can utilize each cell and channel available in the rotor for higher throughput. Like SV experiments, ABDE experiments also benefit from multi-wavelength analysis and can distinguish between DNA and protein content in each species. Examples are shown in Figs. 6-8:
Acknowledgments:
We thank Xiaozhe Ding and Viviana Gradinaru (Caltech), and Oka Kazuhiro and Austin Seal (Baylor College of Medicine) for providing samples for this study. The work was funded by NIH, NSERC, CFI, and the Canada Research Chairs program.
References:
[1] Demeler, B Cur. Protoc. Prot. Sci. (2010) Chapter 7:Unit 7.13. DOI: 10.1002/0471140864.ps0713s60
[2] Demeler B., and G. Gorbet. Ch. 8, Eds: Uchiyama S., Stafford W. F. and T. Laue. Springer, 119-143 (2016)
[3] Berkowitz SA, Laue T. Anal Biochem. 2021 Oct 15;631:114306. doi: 10.1016/j.ab.2021.114306.
[4] Savelyev A, Gorbet GE, Henrickson A, Demeler B. PLoS Comput Biol. 2020 Jun 19;16(6)
Figure 6:
Full AAV capsids measured in MW-AUC/ABDE experiments. Left: Multi-wavelength plot of last scan, Right: Deconvoluted protein (blue) and DNA (red).
Figure 7:
Heterogeneous AAV capsids measured in MW-AUC/ABDE experiments. Left: Multi-wavelength SV results. Right: Same sample analyzed with MW-ABDE. Empty capsid shows mostly protein signal
In both plots protein = blue, DNA=red
  Filled
Capsids
Partially filled
Capsids
Mostly empty
Capsids
Figure 8:
Multi-wavelength ABDE experiment of a mixture of empty (~6.25 cm) and filled (~ 6.57 cm) capsids, only the last scan is shown. Protein (blue) and DNA (red) are deconvoluted and the empty capsid is clearly visible on the left with only a protein signal having appreciable signal. The full capsids display equal ratio of protein and DNA signal (all data are normalized to relative concentrations).